PRE-MAINTENANCE REPORT FOR THE STATIONS LOCATED IN RWAMAGANA, KAYONZA, AND KIREHE DISTRICTS

- INTRODUCTION

This report presents the initial findings on the automatic weather stations located in Rwamagana, Kayonza, and Kirehe Districts. Therefore, technicians are advised to take these observations into account when planning maintenance activities. The report is based on data collected from July 1st, 2025, to October 10th, 2025.

- FINDINGS

  - RWAMAGANA DISTRICT

- RWAMAGANA AWS

- Reporting status: up to date

- Rainfall: Empty cell in web monitor

- Radiation: Vary between 0 W/m2 and 12 W/m2 which is unrealistic.

- Soil Temperature 10cm: Vary between 00C and 49.10C which is unrealistic.

- Missing data: 9 days, 16 hours, 50 minutes (check the details in provided excel file)..

- ARGs

| SN | Station | Last communication | Issue |
| --- | --- | --- | --- |
| 1 | Musha | 06-Oct-2025 | No issue |
| 2 | Nzige | 21-May-2025 | Not Known |

  - KAYONZA DISTRICT

- KAWANGIRE AWS

- Wind direction: Vary between 100 and 1350 which is unrealistic.

- Ten minutes’ s evaporation: varies between 12.4 mm and 13.2 mm which is unrealistic.

- Soil moisture at 20 Cm: Change suddenly, has repeating value over fixed hours every day (04:20 to 15:30) while other hours’ values are null.

- Missing data: 13 days, 2 hours, 40 minutes (check the details in provided excel file).

- MUKARANGE AWS

- Reporting status: up to date

- Rainfall: From 1/07/2025 recorded total rainfall of 0 mm.

- Soil Temperature 10cm: Rise gradually up to 430C.

- Missing data: 11 days, 13 hours, 40 minutes (check the details in provided excel file).

- AKAGERA RWISIRABO AWS

- Reporting status: up to date

- Missing data: 40 days, 17 hours, 0 minutes (check the details in provided excel file).

- NDEGO AWS

- Reporting status: up to date

- Missing data: 17 days, 5 hours, 20 minutes (check the details in provided excel file).

- ARGs

| SN | Station | Last communication | Issue |
| --- | --- | --- | --- |
| 1 | Gahini | 06-Oct-2025 | No issue |
| 2 | Kabarondo | 06-Oct-2025 | No issue |
| 3 | Murundi East | 06-Oct-2025 | No issue |
| 4 | Mwiri | 06-Oct-2025 | Some hours Battery reads 0 V but it reports |

  - KIREHE DISTRICT

- KIREHE AWS

- Reporting status: up to date

- Rain: from 1/07/2025 to 24/09/2025 records total rainfall of 7.2 mm this is contrary to nearby manual station which records 15.8mm from 1/07/2025 to 31/08/2025.

- Wind Direction: This is suspected due to data varies between 59 degrees and 274 degrees.

- Missing data: 22 days, 17 hours, 20 minutes (check the details in provided excel file).

- RUSUMO AWS

- Reporting status: up to date

- Missing data: 35 days, 16 hours, 0 minutes; during night hours did not reports (check the details in provided excel file).

- MAHAMA AWS

- Reporting status: up to date

- Rain: from 1/7/2025 to 24/09/2025 records total rainfall of 0mm.

- Soil Temperature 50cm: This is suspected due to data varies between 31.60C degrees and 39.20C.

- Missing data: 29 days, 12 hours, 30 minutes; below 15/09/2025 report did not report during night hours (check the details in provided excel file).

- NASHO_MPANGA AWS

- Reporting status: up to date

- Wind Direction: Always records 0 degrees.

- Soil Temperature 10cm: the data are not available

- Missing data: 12 days, 0 hours, 30 minutes (check the details in provided excel file).

- ARGs

| Station Name | LAST COMMINICATION | ISSUE |
| --- | --- | --- |
| Gahara | 22-Jun | Not known |
| Musaza | 27-Mar | Not known |
| Nyamugali | 23-May | Not known |
| Nyarubuye Parish | 15-Apr | Not known |

Prepared By:

Christian Bikorimana

Ag.Data Quality Control Officer