REPORT ON MONITORING HYDROGEN GENERATOR PERFORMANCE

- Background

Following the installation of the Upper Air Station and the Hydrogen (H₂) Generator at the Huye Upper Air Station site, an operational malfunction was observed whereby the generator failed to automatically fill water into the electrolytic cells. Consequently, a fourteen days monitoring exercise was conducted from 30th January to 12nd February 2026, in collaboration with the supplier, to diagnose the issue.

- Objective

The objective of the monitoring exercise was to assess the operational performance, reliability, and safety of the Hydrogen generator under normal operating conditions.

- Key Findings

- It was found that the auto-filling system failed due to a malfunction of the KA20 timer relay. Initially, the relay was configured to close the auto-fill system circuit in 10 minutes after the generator started; however, this 10-minute setting was faulty. After reconfiguring the timer to 1 minute, the auto-filling system operated normally.

- It was noted that excessive foam was present in the cells, which misled the filling probe and affected the timely filling of water into the cells.

- The generator demonstrated overall good performance in hydrogen production and maintained safety during operations.

- Recommendations

- To restore the original 10 minutes setting for closing the auto-fill system circuit after generator start-up, the KA20 timer relay needs to be replaced.

- The supplier indicated that the foam was likely caused by chemical agents present in the on-site tubing and is expected to dissipate naturally as the generator continues to operate; therefore, close monitoring is recommended during the initial operational period.

## Conclusion

The monitoring exercise confirmed that the Hydrogen generator is capable of automatically refilling water following reconfiguration of the KA20 timer relay; however, replacement with a new KA20 timer relay will ensure the full functioning of the generator. Foam-related interference with the filling probe continues to affect full automation of the water-filling process. Therefore, further close monitoring during the operational period will determine appropriate corrective actions should take if the issue persist.

Prepared by:

BIKORIMANA Christian                                                                NDAYISHIMIYE Eric

Ag. Data Quality Control Officer                                                     Electrical Engineer

Approved By:

Dr. Vedaste IYAKAREMYE

SOFF Focal Person and Division Manager of DOQC&P Division